## Slide 1: Controlling LED-Illumination by NMR-Spectrometer with Arduino and Python
Franz Westermair

## Slide 2: Hardware to be used
Thorlabs Fiber coupled LEDs with LEDD1B driver, can be controlled with CW, PWM (active high), or analog input (0…5V).
300 MHz Bruker spectrometer with 5 mm probe using 1000 µm fiber (1000UMT) SMA connector.
Arduino Uno R3 or Mega2560 to generate the PWM Signal, need 2-Wire/BNC to BNC cable to connect to LED driver
Spectrometer computer which is capable of running Python 3.4+ for serial communication

## Slide 3: Software requirements
Arduino IDE to upload program to Arduino UNO R3 or MEGA2560
Code works for both Boards without modification
Need to have Python 3 installed
The serial communication is done with the "PySerial" module
TopSpin with Jython of any version as long as it has "os" module

## Slide 4: Setting up the controller (2)
In the script "wef_arduino_sys_py.py", change the COMPORT to what you have determined and check if BAUDRATEis the same as in the main.ino file that was uploaded.
The function can be checked now by running the script in a terminal, adding an integer number (<256) as argument.
You should see "Received: 0:<the entered number>
With higher baudrates it may also print some more, but works fine

## Slide 5: Setting up the controller: TopSpin (3)
copy wef_arduino_ts_py into the "/exp/stan/nmr/py/user" directory
Change the variables pythonpath and scriptpath accordingly
Here is also the place to program which commands, values to send since you have access to all TS parameters here
Now the "wef_arduino_ts_py " script can be called in AU programs via XPY("wef_arduino_ts_py.py") command
There might be a more direct way to do this, but it works fine...
In the " wef_arduino_ts_py" script, you can specify what should be sent as arguments to the controller

## Slide 6: Arduino Standard UART Specification
Baudrate can be set arbitrarily, but recommended to use 9600 or 115200.
Data bits: 8
Parity: None
Stop bits: 1
Flow control: None
(Can be changed by flipping some register bits)

## Slide 7: TopSpin AU programming
Mixture of many Bruker-defined C-Macros defined either in <tshome>/prog/include/aucmd.h //Standard commands or in the AU-directory <tshome>/exp/stan/nmr/au/src